www.wunder.it

PROFESSIONAL ELECTRONIC MDD SCALE

MOD. RH Handrail

| USER MANUAL |
| --- |
| Read this manual carefully before using the instrument |
|  | INDEX |  |

2. GENERAL RULES	2

3. SAFETY	3

4. THECHNICAL SPECIFICATION	6

5. INDICATOR	6

6. INSTALLATION	8

7. INSTRUCTION FOR USE	10

8. PRINTER FUNCTIONS	12

9. SETTING SETUP	14

10. ERROR MESSAGES	14

11. MAINTENANCE AND ASSISTANCE	15

12. CONFORMITY	15

13.IDENTIFICATION LABELS	15

14. SCRAPPING AND WASTE DISPOSAL	16

15. WARRANTY	17

16. ASSEMBLY-USE MECHANICAL TELESCOPIC HEIGHT METER (OPTIONAL)	18

17. THERMAL PRINTER (OPTIONAL)	20

By choosing the WUNDER mod. RH handrail professional electronic scale, you have purchased a high precision instrument. Since over 40 years Wunder has placed its experience at the service of health.  This instrument is compliant with national standards in hospitals and clinics with medical class I with measurement function and is calibrated in conformity with accuracy class III. The instrument is characterised by the possibility of fixing the electronic terminal to the weighing platform by means of a column or else installing the terminal autonomously.

The instrument is equipped with a dual LCD electronic terminal with triple reading to simultaneously view Weight, Height. and BMI

# 2. GENERAL RULES

Carefully read this manual before using the instrument as it supplies important indications concerning OPERATING SAFETY AND MAINTENANCE.

The descriptions and illustrations supplied in this manual are not binding.

WUNDER reserves the right to modify its products as deemed convenient in order to approve them without being committed to update this publication.

Conventions: The following symbols have been used in this manual:

|  | 0476 EC TYPE APROVAL FOR MEDICAL USE |
| --- | --- |
|  | METROLOGICAL CERTIFICATION AND APPROVAL |
|  | ACCURACY CLASS |
|  | ATTENTION! PLACED BEFORE DETERMINING PROCEDURES. COMPLIANCE FAILURE CAN HARM THE OPERATOR OR PATIENT OR DAMAGE THE PRODUCT |
|  | WASTE DISPOSAL |
|  | TYPE B PARTS SUPPLIED |
|  | BATTERY POWER |
|  | INDICATION OF WEIGHT FUNCTIONALITY |
| (O( | INDICATION OF STABLE WEIGHT |
|  | INTERFERENCE COULD OCCUR NEAR THE APPLIANCES |
|  | DUAL INSULATION (CLASS II) |

# 3. SAFETY

ATTENTION!

Operators must read this manual carefully, comply with the instructions it contains and become familiar with the correct use and maintenance procedures of the instrument.

- Retain this manual for consultation and as a help in staff training

- Do not overload the instrument beyond its maximum capacity

- Do not apply loads abruptly.

- Do not press the keys with sharp or pointed objects

- Do not try to open the instrument.

- Do not remove seals from the instrument.

- Do not short-circuit the battery terminals

- Use only the power supply provided by Wunder. Before using it, make sure that the local mains voltage is compatible with the voltage of the adapter shown on the identification plate

- Regularly check the integrity of the instrument's power cord and make sure it does not come in contact with hot appliances

- Make sure that the power cord does not create obstruction hazards

- Unplug the instrument before cleaning it

- Do not place the instrument in water or other liquids

- Perform maintenance and subsequent metric verifications regularly

3.1 Manufacturer's guide and declaration – electromagnetic immunity

| Guide and Statement of manufacturer – Electromagnetic emissions |
| --- |
| The electronic scales RH model is scheduled for operation in the electromagnetic environment specified below. The customer and the user should ensure that it is used in that environment |
| Emission test | Conformity | Guide to electromagnetic environment |
| RF Emission CISPR11 | Group 1 Class B | The scales RH model uses RF energy only for its internal function. The refore its RF emissions are very low and probably not cause no interference in electrical equipment. |
| RF Emission CISPR11 | Group 1, Class B | The scales RH is suitable for use in all sanitary and hospital buildings, connected to the low voltage public power supply network. |
| Harmonic emission | Class A |  |
| Voltage fluctuations/ flicker emission | Compliant |  |

| Guide and Statement of manufacturer-Electromagnetic Immunity |
| --- |
| The electronic scales RH model is scheduled for operation in the electromagnetic environment specified below. The customer and the user should ensure that it is used in that environment. |
| Immunity test | Compliance | Guide to electromagnetic environment |
| Electrostatic discharges (ESD) IEC/EN61000 - 4 -2 | 6kV in contact 8kV in air | The floors should be made of wood, concrete or ceramic. If the floors are covered in synthetic material, the relative humidity should be maximum 30%. |
| Rapid transients/burst IEC/EN61000 - 4 - 4 | 2kV power supply | The power supply should be of the type used typically in commercial or hospital environments. |
| Surge IEC/EN61000 - 4 - 5 | 1kV differential mode | The power supply should be of the type used typically in commercial or hospital environments. |
| Voltage dips, short interruption and voltage variation IEC/EN61000 - 4 - 11 | <5%UT for 0.5 cycle 40%UT for 05 cycle 70%UT for 25 cycle <5%UT for 5 sec | The power supply should be of the type used typically in commercial or hospital environments. Note=UT is the value of the voltage of the feed. |
| Power frequency Magnetic field IEC/EN61000 - 4 - 8 | 3A/m |  |

NOTE: UT is the value of the power supply voltage.

CAUTION

The medical device requires particular electromagnetic compatibility precautions and must be installed and used according to the information provided in the accompanying documents.

| Manufacturer's guide and declaration - Electromagnetic emissions |
| --- |
| The electronic scales RH model can be used in the electromagnetic environment specified below. The customer or the user of the seat must ensure that it is used in compliance with the specifications below. |
| Emissions test | Compliance | Guide to electromagnetic environment |
| Radiated/ conducted / emissions IEC/EN61000-4-6 | 3Vrms 150kHz to 80MHz (for appliances that are not life supporting) | Equipment in RF communications and mobile laptops should ot be used near any part of the tooth, including cables, except when they meet the recommended separation distance calculated by the frequency applied to the transmitter. Recommended separation distance d = 1.2 √ P d = 1.2 √ P from 80 MHz to 800 MHz d = 2.3 √ P from 800 MHz to 2.5 GHz where the P and maximum rated power output of the transmitter in Watts (W) according to the manufacturer of the transmitter and the recommended separation distance in meters (m). The intensity of the field of a fixed RF transmitter, as determined by an electromagnetic site, it could be lower than the level of compliance of each range of frequency b. You can check interference in proximity of equipment marked with the following symbol: |
| Radiated/ conducted /emissions IEC/EN61000-4-3 | 3Vrms 150kHz to 80MHz(for appliances that are not life equipment) |  |
| With 80 MHz and 800 MHz is applied the higher frequency range. These guidelines may not apply in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects and people. a) The intensity of the field for fixed transmitters such as base stations for radiotelephone mobile and cordless phones) and land radio mobile, amateur radio equipment, radio transmitters in the AM and FM and TV transmitters cannot be made theoretically and with precision. To establish an electromagnetic environment due to fixed RF transmitters, it should consider the electromagnetic survey of the site. If the field strength measured at the place where you use the instrument exceeds the applicable level of compliance of the above, it should be placed under observation normal operation of the tooth device. If you notice abnormal performance, it may take additional measures as a different orientation or position of instrument. of frequencies from 150 kHz to 80 MHz should be less than 3 V / m. |

| Recommended separation distances between portable radio equipment unity. |
| --- |
| The scales RH model is expected to operate in an electromagnetic environment in which they control the noise radiated RF. The customer or the tooth may be of help to prevent electromagnetic interference by ensuring a minimum distance between mobile communication devices and portable RF (transmitters) and unity, as recommended below, with respect to the maximum power output of equipment radio. |
| Output power rating of the transmitter W | Distance separating the frequency of the transmitter m |
|  | 150 kHz a 80 MHz d = 1,2 √P | 80 MHz a 800 MHz d = 1,2 √P | 800 MHz a 2,5 GHz d = 2,3 √P |
| 0,01 | 0,12 | 0,12 | 0,23 |
| 0,1 | 0,38 | 0,38 | 0,73 |
| 1 | 1,2 | 1,2 | 2,3 |
| 10 | 3,8 | 3,8 | 7,3 |
| 100 | 12 | 12 | 23 |
| For transmitters with maximum rated power output not reported above, the recommended separation distance d in meters (m) can be calculated using the equation applicable to the frequency of the transmitter, where P and the maximum rated power output of the transmitter in Watt (W) according to the manufacturer of the transmitter. Note: With 80 MHz and 800 MHz is applied the higher frequency range. These guidelines may not apply in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects and people. |

# 4. THECHNICAL SPECIFICATION

| Model | RH 300 |  |
| --- | --- | --- |
| Capacity | 300 kg |  |
| Division | 50- 100 g |  |
| MDD OIML | Class III |
| Unit weight | kg |
| Display | Double LCD: 1 LCD 20mm weight with 5 active digits; 2 LCD height and BMI 12mm |
| Function Key | ON / ZERO / OFF, HOLD / BMI, TARA, PRINT |
| Power | Rechargeable Batteries Pack- Adapter 12 V - 2 A |
| Operation temperature | da 5°C a 35°C |

# 5. INDICATOR

|  | 1. Key function 2. Weighing capacity and division 3. Second LCD automatically calculates BMI 4. Main LCD: Weight 5. Power display |
| --- | --- |

Function Keys

| Key | Description |
| --- | --- |
|  | On/Off key. Pressed for 3 seconds it switches off the scale. |
|  | Indication reset (within ±2% of the capacity) |
|  | Lock of displayed weight/determination of the stable weight value. |
|  | Determine the body mass index (BMI) |
|  | Tare of unwanted weight. |
|  | To delete the incorrect entry during the entering of the digits. |
|  | To confirm the functions. |
|  | Print of data. |
| 0-9 | Digits entering keys. |

Display Functions

Weight stable: Indicates the weight is stable.

Weight negative: Indicates weight below zero.

Zero: Indicates zero weight.

Flat battery: Indicates the need to recharge or replace the battery.

# 6. INSTALLATION

ATTENTION!

AFTER HAVING VERIFIED THE INTEGRITY OF THE EQUIPMENT AND OF ALL ITS COMPONENTS, CONNECT THE INSTRUMENT WITH THE SUPPLIED POWER SUPPLY TO CORRECTLY RECHARGE THE BATTERIES

Position the instrument on a level and stable surface

The instrument must be placed away from heat sources, in an environment without excessive vibrations and drafts.

Assembly

| STAGES TO ASSEMBLY 1. ASSEMBLY DISPLAY 2. ASSEMBLY COLUMN 3. ASSEMBLY LATERAL HANDRAIL 4. ASSEMBLY CIRCULAR HANDRAIL |  |  |
| --- | --- | --- |

| 1: ASSEMBLY DISPLAY |
| --- |
| 1 ( | 1A |
| Take the column put the cable display into the central hole of the plastic support mounted on the column | Fix the display with the screws (m4x0,7*10mm) to the column support |
| 2 ASSEMBLY COLUMN |
| 2 ( | 2A |
| Take white metal angular and insert the connect cable through the center hole and fix it to the column with the screws. position it in the corner of platform and connect the cables | Fix the angular to the platform with the screws |
| 3 ASSEMBLY LATERAL HANDRAIL |
| 3 | 3A |
| Fix lateral handrail with supplied screws | Fix2° lateral handrail with supplied screws |
| 4 ASSEMBLY CIRCULAR HANDRAIL |
| 4 ( | 4A |
| Insert circular handrail into lateral and fix it with the screw | Fix the handrail with the supplied screws |
| ( |  |

# 7. INSTRUCTION FOR USE

ATTENTION!

For proper charging of the battery pack, connect the instrument to a power outlet for at least 8 hours

7.1 Instruction for charging and connecting

If  prompt displays on the LCD, please charge the scale with exclusive adaptor.

Locate adaptor plug on the rear side of indicator.

The battery should be recharged at least every 3 months regardless if it is used or not.  After a long period in storage, e.g. over 3 months, the battery should run a full cycle (charge/discharge) to allow it to restore to full capacity.

| 7.1a Rechargeable Batteries |
| --- |
| Open the battery cover and take the battery out. | Disconnect the rechargeable battery pack then replace it by a new one. Put the battery back then arrange the connector in place well. |

7.2 DATE AND TIME SETTING

Keep the HOLD key pressed for 3 seconds to access TIME SETTING programming.

Example: to program 25 December 2016, 8:00 a.m.

|  | Year programming: use the number keys to enter the correct value, press HOLD to access the next step once correctly set. |
| --- | --- |
|  | Date programming: Use the number keys to enter the correct value, press HOLD to access the next step once correctly set. |
|  | Time programming: Use the number keys to enter the correct value, press HOLD to confirm the entered date and time once correctly set. |
|  | Display format: YYYY → MM.DD → HH:MM |

7.3 HOLD FUNCTION

The Hold function is used to accurately determine the average weight, also with unstable value (small movements)

- switch-on the scale using the [ON/OFF] key. The display shows “0.0 kg” and the stable weight and zero indications, after the self-test.

- Ask the patient to climb onto the weighing platform.

- Press HOLD. The display shows "HOLD" with a flashing triangle and then, after a few seconds, the average weight value that remains permanently on.

- Press HOLD again to release the weight memorised on the display. The instrument goes back to normal weighing mode.

- The HOLD function may be activated before or after the patient has climbed onto the platform.

7.4 BMI FUNCTION

| Press BMI in weighing mode. The secondary display shows the last set height, with the left digit flashing. |  |
| --- | --- |
| Use the number keypad to enter the height (e.g.: 170 cm) Confirm with |  |
| Press HOLD to lock the value. Climb onto the weighing platform when the display shows "HOLD" and weigh yourself. The scale shows the weight, height and BMI index. |  |
| Press BMI again to release the memorised weight and go back to normal operation mode. |  |

7.5 TARE FUNCTION

This function allows displaying two weights. Switch-on the instrument using the key  .

When the display indicates  and a beep is emitted, position the objects (clothing) constituting the tare, onto the weighing platform.

With the weight stable and the appearance of the stable symbol (circle), press  and the display shows  . Ask the patient to climb onto the platform. The indicated weight represents the tare's net value. Delete the memorised tare value by climbing down from the instrument and press  again.

# 8. PRINTER FUNCTIONS

Weight can be printed for records using RS232 interface cable.After weighing simply press PRINT key to print out the results. The format presented below is the standard format of results print-out and cannot be changed. More informations  info@wunder.it or service@wunder.it

|  | Set parameters of the scale interface on the connected device. It is not possible to change the scale parameters. Baud rate: 9600 bps Parity check: None Data length: 8 bits Stop bit: 1 bit Handshake: RTS/CTS Data code: ASCII |
| --- | --- |

Serial RS232

| Piedino | Segnale |
| --- | --- |
| 2 | TXD |
| 3 | RXD |
| 5 | GND |

Connecting with PC

Start Hyper Terminal

Start Hyper Terminal program from clicking

Start Menu → Programs → Accessories → Communication → Hyper Terminal.

New Connection Description

Give new connection a name then click OK.

Select Your COM Port

Click Connect to select your COM port. Usually there’s only one option for select. Then click OK.

Port Settings

Click Bits per second to set up rate at 9600, Data bits at 8, Parity at None, Stop bits at 1 and Flow control at Hardware. Then click OK to complete your setting.

Output Data

When press the PRINT key to output data from scale to PC or an Optional Printer.

# 9. SETTING SETUP

Switch ON the scale and long press  key for 3 seconds, first “SETUP” and then “A.OFF” will display successively.

AUTO-OFF TIME SETUP This enables operator to select the auto turn OFF time of the device.

BUZZER ON/OFF SETUP This enables operator to select beep sound ON/OFF.

Note: To confirm the settings, please press  when displays on the display.

# 10. ERROR MESSAGES

# 11. MAINTENANCE AND ASSISTANCE

For better and longer duration of the product it should receive thorough general cleaning periodically.

The instrument must be cleaned with a soft cloth moistened with water or neutral detergent, without using solvents or abrasive substances. If the instrument remains idle for a long period, remove the batteries from the terminal. During shipping, make sure not to subject the instrument to blows or excessive mechanical stress. In case of repairs or assistance, contact your dealer or an authorised service centre contacting service2@wunder.it or sales@wunder.it.

# 12. CONFORMITY

WUNDER MODEL RH or RO CABLE ELECTRONIC SCALE SERIAL N°...............................................................

We hereby certify that this instrument has been inspected and has successfully passed the functional test. Itcomplies with the following standards and directives:

EN 45501 / EN60601-1-2 / EN60601-1 93/42/EEC – Medical Devices Directive

# 13.IDENTIFICATION LABELS

# 14. SCRAPPING AND WASTE DISPOSAL

If set aside for a long period, protect those parts which could be damaged due to dust build-up

Scrapping

When you decide to no longer use this item, we recommend making it unusable. We also recommend making those parts which could be sources of danger harmless

Waste disposal 2012/19/UE

This product complies with the disposal 2012/19/UE. The symbol of the crossed-out waste bin on the appliance indicates that the product, needing to be treated separately from household waste, at the end of its useful life must be completed in a separate collection facility for electric and electronic appliances or returned to the dealer upon purchase of a new equivalent appliance. The user is responsible for bringing the appliance to an appropriate collection structure at the end of its life. Appropriate separate collection and sending the appliance for recycling, treatment and environmentally compatible waste disposal contributes to avoid possible negative effects on the environment and health and favours the recycling of the materials the product is made of.

For more detailed information regarding available collection systems, contact your local waste disposal service or the shop where the product was purchased.

As consumers, you are obliged by law to return used or dead batteries. You may deposit old batteries at public collection spots in your town or else with any battery dealer who has placed specific collectors for this purpose. Even when scrapping electric and electronic appliances, they must be removed and deposited in specific collectors.

NOTE: The following symbols indicate the presence of harmful substances

| Pb Pb = batteries containing Lead | Cd Cd = batteries containing Cadmium | Hg Hg = Batteries containing Mercury |
| --- | --- | --- |

ATTENTION!

Do not throw electric parts and used batteries away with household waste. Dispose of the batteries by means of your closest collection centres.

# 15. WARRANTY

This certificate must be kept until the warranty has expired.

It must be presented together with the invoice, tax receipt or delivery note providing the name of the dealer and date of purchase whenever a technical intervention is required. Otherwise the user will lose any warranty rights.

The warranty takes effect from the date of purchase and is valid during the entire period foreseen by the current catalogue/pricelist.

By warranty we mean the replacement or repair free of charge of parts making up the appliance which, at the discretion of the manufacturer, are deemed faulty from the origin; Wunder therefore has the faculty of repairing or replacing the item.

The warranty does not cover:

- shipping faults, damage caused by falls, carelessness or tampering

- damage caused by incapability of using the appliance and of its improper use

- damage caused by an insufficient or inadequate electrical system or alterations resulting from environmental, climatic or other types of conditions

- damage due to incorrect installation of the appliance and repairs carried out by unauthorised personnel

- Interventions at home for convenience controls or presumed defects

- Routine maintenance and that which can be considered normal wear from use

- consumables such as: power supplies, batteries, keyboards, plates, wheels, heads, rolls, load cells faulty due to blows or overloads

Service can also be refused when the appliance has been changed or transformed in any way.

In case of interventions at one's home, the customer must pay the fixed fee; if however the appliance is repaired at an authorised Wunder Service Centre, expenses and relative travel risk are borne by the user.

Wunder will not be held liable for damage of any nature caused directly or indirectly to persons, animals or objects resulting from failure to comply with all the instructions indicated in this manual or anyway resulting from improper use.

The Court of Bergamo has jurisdiction in case of any di

# 16. ASSEMBLY-USE MECHANICAL TELESCOPIC HEIGHT METER (OPTIONAL)

ATTENTION!

17.1. STEP 1: FIXING MECHANICAL OR ELECTRONIC STATIMETER

| Figure 1a: the instrument in the packaging is presented with the support statimeter already mounted (screwed screw in the holder). It is only necessary to fix the statimeter to the black ABS support Figure 1b: statometer fix the measurement height to the support in black abs with the supplied screw. |
| --- |
| Fig. 1a |  | Fig. 1b |

17.2: USE OF MANUAL MECHANICAL STATIMETER

ATTENTION!

FOR THE MEASUREMENT OF HEIGHT THE PATIENT MUST BE ON THE SCALE PLATFORM

| Misure superiori a 131 cm | Misure inferiori a 130 cm |
| --- | --- |
| 1 | Pull up the height meter rod | 1 | Lower the headrest palette and press the black center stopper |
| 2 | Position it horizontally | 2 | Holding down the black stop, lower the headrest palette |
| 3 | Lower the rod up to the point where the lever rests head touches the patient's head. The height can be read at the point 'A' | 3 | Lower the rod to the point where the blade touches the patient's head. The height can be read at point 'B'. |

16.1. ASSEMBE ELECTRONIC HEIGHT METER (OPTIONAL)

STAGE 1: INSTALLATION ALKALINE BATTERY

|  | Removing battery cover at the rear of indicator Installing code AAA batteries * 2 Reinstalling the battery cover |
| --- | --- |

STAGE 2: CONNECTING CABLE DIGITAL TELESCOPIC HEIGHT METER

The indicator connection is under indicator

The height meter connection is under indicator height meter

STAGE 3: MEASURING HEIGHT

ATTENTION!

THE MEASUREMENT OF THE PATIENT MUST BE GETTING ON THE PLATFORM OF SCALE

16.2 USE DIGITAL TELESCOPIC HEIGHT ROD

Make indicator move to the end of measuring rod. Then switch ON and press reset key once to calibrate.

Steps to accurate measurement

- Push the head stop far enough up the rod so that the person to be measured can stand underneath it comfortably.

- Ask the person to be measured to stand with his or her back to the rod.

- Sliding the head stop down until it touches the head.

- Read the height measured on the display unit.

17. THERMAL PRINTER (OPTIONAL)

Read this manual carefully before using the instrument

17.1. HOW TO USE YOUR PRINTER

Connect the AC Adaptor and RS232 interface cable.  Press the I/O button to power on the printer.

The PRINTER can be connected via the RS232 port and cable with the RB200

Press the PRINT key: Print double language

CHANGING THE PAPER ROLL

The paper print roll on a PRINTER XL printer is changed as follows:

| 1. Open the upper cover and position the paper roll so that it feeds in the right direction. | 2.Feed the end of the paper roll through the printer cover and close the cover. The printer is ready to print. |
| --- | --- |

17.2. PRINTER SPECIFICATION

| ITEM | SPECIFICATION |
| --- | --- |
| Printing method | Thermal dot line printing |
| Number of dots | 384 dot/line |
| Resolution (dots/mm) | 8 dot/mm |
| Printing width (mm) | 48 mm |
| Max. paper thickness (μm) | 80μm |
| Paper | 57.5±0.5 mm (W) x ψ30mm max |
| Paper loading | Easy loading |
| Printing speed (mm/s) | 75 mm/s |
| Dimensions (L x W x H ) | 120 x 80 x 50mm |
| Operating temperature (℃) | 0℃ to 50℃ |
| Operating humidity (%) | 20% to 85% (no condensation) |
| Interface | Conform to RS232 |
| Voltage range (adaptor only) | 12V 2A |

SHOCK HAZARD WARNING

The printer and power supply should never be operated in a location where either one can get wet.  Personal injury could result.

MEDIA WARNING

Always use high quality approved paper print roll.  The approved product can be ordered from your dealer.

RELOADING HINT

If printer runs out of paper roll while printing, switch the power off and reload paper roll.  Power on and print again.

17.3. ERROR MESSAGES

| Problem | Solution or Reason |
| --- | --- |
| Status Indicator does not light when power switch is turned to ON (I) position. | Check the power connections from AC outlet to the printer power supply. |
| Status indicator illuminated (green light), but the printer will not print. | Check interface cable connections from the scale to printer. Make sure the top cover is closed and locked. |
| Printer paper jammed. | Reset the paper roll. Please refer to “CHANGING THE PAPER ROLL”. |
| Status Indicator is flashing green. | Paper roll is out. Reload a new source of paper roll so the printer can continue printing. Please refer to “CHANGING THE PAPER ROLL”. |

# Text boxes

1

4

5

2

3

Weight stable

Weight negative

Zero

Flat battery

kg

Net weight

Weight

Hold Function

Platform

Column

Lateral handrail

Display

Circular handrail

Cable display

IMAGE AFTER THE ASSEMBLY

Ac Jack connector

Connector RS232

Plug power indicator

Plug power digital telescopic height meter

Height meter connector

Indicator connector

Platen Rollen

Tear Bar

Printer Button

Status Indicator

Roll Holder

RS 232 connect

ON/OFF

Power

(Adattatore)