www.wunder.it

WEIGHING PLATFORM WHEELCHAIR

MOD. RW02 - WU150 18 keys

INSTRUCTION MANUAL

Read this manual carefully before using the instrument

INDEX

2. GENERAL RULES	2

3. SAFETY	3

4. TECHNICAL SPECIFICATION	6

5. INDICATOR	6

6. INSTALLATION	7

7. INSTRUCTION FOR USE	9

8. PRINTING FUNCTION	13

9. SETTING SETUP	14

10. ERROR MESSAGE	14

11. MAINTENANCE AND ASSISTANCE	15

12. CONFORMITY	15

13. IDENTIFICATION LABELS	15

14. SCRAPPING AND WASTE DISPOSAL	16

15. WARRANTY	17

By choosing the WUNDER mod. RW (WU150) professional electronic scale, you have purchased a high precision instrument.

Since over 40 years Wunder has placed its experience at the service of health.  This instrument is compliant with national standards in hospitals and clinics with medical class I with measurement function and is calibrated in conformity with accuracy class III.

The instrument is characterised by the possibility of fixing the electronic terminal to the weighing platform by means of a column or else installing the terminal autonomously.

The instrument is equipped with a dual LCD electronic terminal with triple reading to simultaneously view Weight, Height. and BMI

# 2. GENERAL RULES

Carefully read this manual before using the instrument as it supplies

important indications concerning OPERATING SAFETY AND MAINTENANCE.

The descriptions and illustrations supplied in this manual are not binding.

WUNDER reserves the right to modify its products as deemed convenient in order to approve them without being committed to update this publication.

Conventions:

The following symbols have been used in this manual:

|  | 0476 EC TYPE APROVAL FOR MEDICAL USE |
| --- | --- |
|  | METROLOGICAL CERTIFICATION AND APPROVAL |
|  | ACCURACY CLASS |
|  | ATTENTION! PLACED BEFORE DETERMINING PROCEDURES. COMPLIANCE FAILURE CAN HARM THE OPERATOR OR PATIENT OR DAMAGE THE PRODUCT |
|  | WASTE DISPOSAL |
|  | TYPE B PARTS SUPPLIED |
|  | BATTERY POWER |
|  | INDICATION OF WEIGHT FUNCTIONALITY |
| (O( | INDICATION OF STABLE WEIGHT |
|  | INTERFERENCE COULD OCCUR NEAR THE APPLIANCES |
|  | DUAL INSULATION (CLASS II) |

# 3. SAFETY

ATTENTION!

Operators must read this manual carefully, comply with the instructions it contains and become familiar with the correct use and maintenance procedures of the instrument.

- Retain this manual for consultation and as a help in staff training

- Do not overload the instrument beyond its maximum capacity

- Do not apply loads abruptly.

- Do not press the keys with sharp or pointed objects

- Do not try to open the instrument.

- Do not remove seals from the instrument.

- Do not short-circuit the battery terminals

- Use only the power supply provided by Wunder. Before using it, make sure that the local mains voltage is compatible with the voltage of the adapter shown on the identification plate

- Regularly check the integrity of the instrument's power cord and make sure it does not come in contact with hot appliances

- Make sure that the power cord does not create obstruction hazards

- Unplug the instrument before cleaning it

- Do not place the instrument in water or other liquids

- Perform maintenance and subsequent metric verifications regularly

3.1 Manufacturer's guide and declaration – electromagnetic immunity

| Guide and Statement of manufacturer – Electromagnetic emissions |
| --- |
| The electronic scales RW, model is scheduled for operation in the electromagnetic environment specified below. The customer and the user should ensure that it is used in that environment |
| Emission test | Conformity | Guide to electromagnetic environment |
| RF Emission CISPR11 | Group 1 Class B | The scales RW model uses RF energy only for its internal function. The refore its RF emissions are very low and probably not cause no interference in electrical equipment. |
| RF Emission CISPR11 | Group 1, Class B | The scales RW is suitable for use in all sanitary and hospital buildings, connected to the low voltage public power supply network. |
| Harmonic emission | Class A |  |
| Voltage fluctuations/ flicker emission | Compliant |  |

| Guide and Statement of manufacturer-Electromagnetic Immunity |
| --- |
| The electronic scales RW model is scheduled for operation in the electromagnetic environment specified below. The customer and the user should ensure that it is used in that environment. |
| Immunity test | Compliance | Guide to electromagnetic environment |
| Electrostatic discharges (ESD) IEC/EN61000 - 4 -2 | 6kV in contact 8kV in air | The floors should be made of wood, concrete or ceramic. If the floors are covered in synthetic material, the relative humidity should be maximum 30%. |
| Rapid transients/burst IEC/EN61000 - 4 - 4 | 2kV power supply | The power supply should be of the type used typically in commercial or hospital environments. |
| Surge IEC/EN61000 - 4 - 5 | 1kV differential mode | The power supply should be of the type used typically in commercial or hospital environments. |
| Voltage dips, short interruption and voltage variation IEC/EN61000 - 4 - 11 | <5%UT for 0.5 cycle 40%UT for 05 cycle 70%UT for 25 cycle <5%UT for 5 sec | The power supply should be of the type used typically in commercial or hospital environments. Note=UT is the value of the voltage of the feed. |
| Power frequency Magnetic field IEC/EN61000 - 4 - 8 | 3A/m |  |

NOTE: UT is the value of the power supply voltage.

CAUTION

The medical device requires particular electromagnetic compatibility precautions and must be installed and used according to the information provided in the accompanying documents.

| Manufacturer's guide and declaration - Electromagnetic emissions |
| --- |
| The electronic scales RW model can be used in the electromagnetic environment specified below. The customer or the user of the seat must ensure that it is used in compliance with the specifications below. |
| Emissions test | Compliance | Guide to electromagnetic environment |
| Radiated/ conducted / emissions IEC/EN61000-4-6 | 3Vrms 150kHz to 80MHz (for appliances that are not life supporting) | Equipment in RF communications and mobile laptops should ot be used near any part of the tooth, including cables, except when they meet the recommended separation distance calculated by the frequency applied to the transmitter. Recommended separation distance d = 1.2 √ P d = 1.2 √ P from 80 MHz to 800 MHz d = 2.3 √ P from 800 MHz to 2.5 GHz where the P and maximum rated power output of the transmitter in Watts (W) according to the manufacturer of the transmitter and the recommended separation distance in meters (m). The intensity of the field of a fixed RF transmitter, as determined by an electromagnetic site, it could be lower than the level of compliance of each range of frequency b. You can check interference in proximity of equipment marked with the following symbol: |
| Radiated/ conducted /emissions IEC/EN61000-4-3 | 3Vrms 150kHz to 80MHz (for appliances that are not life equipment) |  |
| With 80 MHz and 800 MHz is applied the higher frequency range. These guidelines may not apply in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects and people. a) The intensity of the field for fixed transmitters such as base stations for radiotelephone mobile and cordless phones) and land radio mobile, amateur radio equipment, radio transmitters in the AM and FM and TV transmitters cannot be made theoretically and with precision. To establish an electromagnetic environment due to fixed RF transmitters, it should consider the electromagnetic survey of the site. If the field strength measured at the place where you use the instrument exceeds the applicable level of compliance of the above, it should be placed under observation normal operation of the tooth device. If you notice abnormal performance, it may take additional measures as a different orientation or position of instrument. of frequencies from 150 kHz to 80 MHz should be less than 3 V / m. |

| Recommended separation distances between portable radio equipment unity. |
| --- |
| The scale RW model is expected to operate in an electromagnetic environment in which they control the noise radiated RF. The customer or the tooth may be of help to prevent electromagnetic interference by ensuring a minimum distance between mobile communication devices and portable RF (transmitters) and unity, as recommended below, with respect to the maximum power output of equipment radio. |
| Output power rating of the transmitter W | Distance separating the frequency of the transmitter m |
|  | 150 kHz a 80 MHz d = 1,2 √P | 80 MHz a 800 MHz d = 1,2 √P | 800 MHz a 2,5 GHz d = 2,3 √P |
| 0,01 | 0,12 | 0,12 | 0,23 |
| 0,1 | 0,38 | 0,38 | 0,73 |
| 1 | 1,2 | 1,2 | 2,3 |
| 10 | 3,8 | 3,8 | 7,3 |
| 100 | 12 | 12 | 23 |
| For transmitters with maximum rated power output not reported above, the recommended separation distance d in meters (m) can be calculated using the equation applicable to the frequency of the transmitter, where P and the maximum rated power output of the transmitter in Watt (W) according to the manufacturer of the transmitter. Note: With 80 MHz and 800 MHz is applied the higher frequency range. These guidelines may not apply in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects and people. |

# 4. TECHNICAL SPECIFICATION

| Model | RW 02 |
| --- | --- |
| Capacity | 300kg x 100g |
| OIML Approval | Classe III |
| Weighing Unit | Kg |
| LCD Display | Double LCD: 1 LCD 20mm weight with 5 active digits; 2 LCD height and BMI 12mm |
| Key Functions | ON/ZERO/OFF, HOLD/BMI, TARA, PRINT, MEMORY |
| Power Supply | Adapter 12 V - 2 A Rechargeable Batteries Pack- (OPTIONAL) |
| Operation Temperature | 5°C / 35°C |

# 5. INDICATOR

|  | 1. Key function 2. Weighing capacity and division 3. Second LCD automatically calculates BMI 4. Main LCD: Weight 5.Power display |
| --- | --- |

Function Keys

| Key | Description |
| --- | --- |
|  | On/Off key. Pressed for 3 seconds it switches off the scale. |
|  | Indication reset (within ±2% of the capacity) |
|  | Lock of displayed weight/determination of the stable weight value. |
|  | Determine the body mass index (BMI) |
|  | Tare of unwanted weight. |
|  | To delete the incorrect entry during the entering of the digits. |
|  | To confirm the functions. |
|  | Print of data. |
| 0-9 | Digits entering keys. |

Display Functions

Weight stable: Indicates the weight is stable.

Weight negative: Indicates weight below zero.

Zero: Indicates zero weight.

Flat battery: Indicates the need to recharge

# 6. INSTALLATION

After unpacking the unit, check the integrity and the submission of all components.
The instrument should be situated away from heat sources, in an environment free from excessive vibration and air currents. The weighing platform must be placed on a flat, stable surface, and all the feet should rest on the ground.

Version standard with cable:

The instrument is ready for use

6.1 ASSEMBLY INSTRUCTION FOR THE OPTIONAL ACCESSORY

WALL MOUNTING BRACKET:

RW (150) Version with cable with wall support (optional included)

1. Place the metal bracket to the wall and properly mark the holes Insert dowels into the holes. Fix the wall support with screw included.

2. 3.

Fix the indicator support with screw included.

STAND FOR THE INDICATOR:

RW (150) Version stand display (optional not included only request)

1. The user can decide to connect the cable inside or outside column.

2. Fix the display with screws included to the stand.

# 7. INSTRUCTION FOR USE

7.1 Time setting

Long press PRINT key for 3 seconds to enter the TIME SETTING mode, beginning with the top row, with the flashing digit.  Press PRINT key after successful change to move to next step.

EX: To input Dec 25, 2016, 8:00am.

|  | Enter year:Using TARE keys to decrease the value, press HOLD key to move to next digit. Press PRINT key after successful entry to move to next stage. |
| --- | --- |
|  | Enter date:Using TARE keys to decrease the value, press HOLD key to move to next digit. Press PRINT key after successful entry to move to next stage. |
|  | Enter time: Using TARE keys to decrease the value, press HOLD key to move to next digit. Press PRINT key after successful setting. |
|  | Display Format YYYY→MM.DD→TT:SS |

ATTENTION!

For proper charging of the battery pack, connect the instrument to a power outlet for at least 8 hours

7.2 Instruction for charging and connecting

If  prompt displays on the LCD, please charge the scale with exclusive adaptor.

Locate adaptor plug on the rear side of indicator.

The battery should be recharged at least every 3 months regardless if it is used or not.  After a long period in storage, e.g. over 3 months, the battery should run a full cycle (charge/discharge) to allow it to restore to full capacity.

7.3. Instruction for replacing batteries

| Open the battery cover and take the battery out. | Disconnect the rechargeable battery pack then replace it by a new one. Put the battery back then arrange the connector in place well. |
| --- | --- |

7.4.  Weighing

Switch on the scale using  key. The diagnostic scale self-check is performed and the software version is displayed.

The „0,00 kg” weight displays on the screen, now the scale is ready for weighing.

Note: If display is not at 0.0kg while switching power on.  Kindly press  key to zero the scale which can be used at any time to zero the

7.5. Hold function

The scale is provided with the integrated hold function to determine the average weight. It enables people to be weighed accurately although they are not still on the scale platform.

- Switch ‘ON’ the scale using [ON/ZERO/OFF] key. The diagnostic self-check will be performed and after that „0.0 kg” will display on the screen with ‘stable’ and ‘zero sign’.

- Move the object/ person on the scale platform.

- Press the [HOLD/BMI] key once. The ‘HOLD’ will display on the screen with blinking triangle and after few seconds the average of fluctuating weight will display on the screen and will lock on the display.

- To release the locked weight on display simply press the [HOLD/BMI] key again and the display will return to normal mode automatically.

- [HOLD/BMI] key function can be activated before or after putting the weight on the tread platform. But in case of weighing unstable person it is recommended to press [HOLD/BMI] key after the person moves on the tread platform.

7.6. BMI function

| In normal mode, long press [HOLD/BMI] key 3sec to enter into BMI mode. The second display will show last height and the extreme left digit will blink. |  |
| --- | --- |
| Use [TARE] to decrease value and press [HOLD] to change to next digit. Press [ON/ZERO/OFF] key to confirm height value. |  |
| When the display shows “Hold”, please stand on the scale platform and proceed to weigh as normal. The instrument will show the weight, height, and BMI value. |  |
| To release the locked weight on display simply press the [HOLD/BMI] key again and the display will return to normal mode automatically. |  |

7.7. Tare function

Tare allows the user to zero the instrument to cancel the weight of a container/ clothes from the reading of the instrument, thus giving the true weight of the product/ person being tested.

Turn on the scale using  key.

When  displays on the screen and scale beeps twice, Place the object (clothes) need to be tare on the tread platform.

Press  after the weight stabilizes and stable sign appears on the display.

Display will return to

Place the item/ baby need to be weighed (without removing the tare object). The weighing result is the Net weight of the item / baby.

To delete the saved tare value, remove the tare object from the tray and press  key again.

7.8. Pre-tare function

The scale can store 5 sets of pre-tare weight.  To use dead weight to store pre-tare value at

m1 (memory 1) ~ m5 (memory 5) as below:

| DESCRIPTION | EXAMPLE |
| --- | --- |
| Press “MEMORY” key after loading the weight on the platform; the display shows blinking “mX” sign (ex. m1) |  |
| Press “TARA” key to address the memory. (m1 ~ m5) |  |
| Press “ZERO” key to store the pre-tare weight; the instrument will be sounding “beep”. |  |

Using Dead Weight:

7.9. Recall Pre-Tare Weight

| DESCRIPTION | EXAMPLE |
| --- | --- |
| Long press “MEMORY” key for 3 seconds; the display will show pre-tare weight of m5 first. The pre-tare weight will blink and a annunciator under “pretare” will be turned on. |  |
| Press “TARA” key till the intended memory of pre-tare weight displayed. |
| Press “ZERO” key to confirm the pre-tare weight to be used; the instrument will automatically perform tare function. |  |
| Press “MEMORY” key, return to Weighing Mode. Pretare annunicator will be turned off. |  |

# 8. PRINTING FUNCTION

Weight can be printed for records using RS232 interface cable.

After weighing simply press PRINT key to print out the results.

The format presented below is the standard format of results print-out and cannot be changed.

Parameters of RS232 interface

Seriale RS232

Connecting with PC

Start Hyper Terminal program from clicking

Start Menu → Programs → Accessories → Communication → Hyper Terminal.

| Pin | Signal |
| --- | --- |
| 2 | TX |
| 3 | RX |
| 5 | GND |

- New Connection Description

Give new connection a name then click OK.

- Select Your COM Port

Click Connect to select your COM port. Usually there’s only one option for select. Then click OK.

- Port Settings

Click Bits per second to set up rate at 9600, Data bits at 8, Parity at None, Stop bits at 1 and Flow control at Hardware. Then click OK to complete your setting.

- Output Data

When press the PRINT key to output data from scale to PC or an

# 9. SETTING SETUP

Switch ON the scale and long press  key for 3 seconds, first “SETUP” and then “A.OFF” will display successively.

AUTO-OFF TIME SETUP: This enables operator to select the auto turn OFF time of the device.

BUZZER ON/OFF SETUP: This enables operator to select beep sound ON/OFF.

Note: To confirm the settings, please press  when displays on the display.

# 10. ERROR MESSAGE

# 11. MAINTENANCE AND ASSISTANCE

Cleaning the instrument must be done with a soft cloth dampened with water or neutral detergent, avoiding the 'use of solvents or abrasives.
During transportation, be careful not to subject the instrument to excessive physical shock or mechanical stress. In case of repair or service, contact your dealer

# 12. CONFORMITY

WEIGHING INSTRUMENT DIGITAL WUNDER: RW (WU150) MODEL SERIAL NUMBER...................

This is to certify that this instrument has been tested and has passed the functional test. Meets the following standards:

EN 45501 / EN60601-1-2 / EN60601-1 93/42/EEC – Medical Devices Directive

# 13. IDENTIFICATION LABELS

# 14. SCRAPPING AND WASTE DISPOSAL

If set aside for a long period, protect those parts which could be damaged due to dust build-up

Scrapping

When you decide to no longer use this item, we recommend making it unusable. We also recommend making those parts which could be sources of danger harmless

Waste disposal 2012/19/UE

This product complies with the EU Directive 2012/19/UE. The symbol of the crossed-out waste bin on the appliance indicates that the product, needing to be treated separately from household waste, at the end of its useful life must be completed in a separate collection facility for electric and electronic appliances or returned to the dealer upon purchase of a new equivalent appliance. The user is responsible for bringing the appliance to an appropriate collection structure at the end of its life. Appropriate separate collection and sending the appliance for recycling, treatment and environmentally compatible waste disposal contributes to avoid possible negative effects on the environment and health and favours the recycling of the materials the product is made of.

For more detailed information regarding available collection systems, contact your local waste disposal service or the shop where the product was purchased.

As consumers, you are obliged by law to return used or dead batteries. You may deposit old batteries at public collection spots in your town or else with any battery dealer who has placed specific collectors for this purpose. Even when scrapping electric and electronic appliances, they must be removed and deposited in specific collectors.

NOTE: The following symbols indicate the presence of harmful substances

| Pb Pb = batteries containing Lead | Cd Cd = batteries containing Cadmium | Hg Hg = Batteries containing Mercury |
| --- | --- | --- |

ATTENTION!

Do not throw electric parts and used batteries away with household waste. Dispose of the batteries by means of your closest collection centres.

# 15. WARRANTY

This certificate must be kept until the warranty has expired. It must be presented together with the invoice, tax receipt or delivery note providing the name of the dealer and date of purchase whenever a technical intervention is required. Otherwise the user will lose any warranty rights.

The warranty takes effect from the date of purchase and is valid during the entire period foreseen by the current catalogue/pricelist.

By warranty we mean the replacement or repair free of charge of parts making up the appliance which, at the discretion of the manufacturer, are deemed faulty from the origin; Wunder therefore has the faculty of repairing or replacing the item.

The warranty does not cover:

- shipping faults, damage caused by falls, carelessness or tampering

- damage caused by incapability of using the appliance and of its improper use

- damage caused by an insufficient or inadequate electrical system or alterations resulting from environmental, climatic or other types of conditions

- damage due to incorrect installation of the appliance and repairs carried out by unauthorised personnel

- Interventions at home for convenience controls or presumed defects

- Routine maintenance and that which can be considered normal wear from use

- consumables such as: power supplies, batteries, keyboards, plates, wheels, heads, rolls, load cells faulty due to blows or overloads

Service can also be refused when the appliance has been changed or transformed in any way.

In case of interventions at one's home, the customer must pay the fixed fee; if however the appliance is repaired at an authorised Wunder Service Centre, expenses and relative travel risk are borne by the user. Wunder will not be held liable for damage of any nature caused directly or indirectly to persons, animals or objects resulting from failure to comply with all the instructions indicated in this manual or anyway resulting from improper use. The Court of Bergamo has jurisdiction in case of any dispute.

# Text boxes

4

5

1

2

3

Weight stable

Weight negative

Zero

Flat battery

kg

Net weight

Weight

HOLD Function

1

9.6cm

Wall pin plug 1”

Self tapping screw

4*20

3

2

Phillips screw

M4 x 0.7 x 10mm

1

M4 x 0.7 x 12

2

AC Jack connector

Connector RS232