www.wunder.it

ELECTRONIC HANGING SICK LIFT SCALE MOD. RS300

| USER MANUAL |
| --- |
|  | Read this manual carefully before using the instrument |  |
|  | INDEX |  |

2. GENERAL RULES	2

3. SAFETY	3

4. FEATURES	7

5. KEY PANNEL	7

6. WEIGHING OPERATION	8

7. FUNCTIONS SETTING	9

8. OPERATION FLOW	9

9. SETTING UP YOUR SCALE STEP BY STEP	10

10. INSTRUCTION FOR BATTERY INSTALLATION	10

11. WARNING INFORMATION	11

12. CLEANING & MAINTENANCE	11

13. DISPOSAL OF THE SCALE	11

14. WARRANTY-LIABILITY	12

15. IDENTIFICATION PLATES	12

By choosing the WUNDER mod.RS300 professional electronic hanging scale, you have purchased a high precision instrument.

Since over 40 years Wunder has placed its experience at the service of health.

This instrument is compliant with national standards in hospitals and clinics with medical class I with measurement function and is calibrated in conformity

with accuracy class III.

The instrument is equipped with large LCD electronic terminal with function Hold weight.

# 2. GENERAL RULES

Carefully read this manual before using the instrument as it supplies

important indications concerning OPERATING SAFETY AND MAINTENANCE.

The descriptions and illustrations supplied in this manual are not binding.

WUNDER reserves the right to modify its products as deemed convenient in order to approve them without being committed to update this publication.

Conventions:

The following symbols have been used in this manual:

|  | 0476 EC TYPE APROVAL FOR MEDICAL USE |
| --- | --- |
|  | METROLOGICAL CERTIFICATION AND APPROVAL |
|  | ACCURACY CLASS |
|  | ATTENTION! PLACED BEFORE DETERMINING PROCEDURES. COMPLIANCE FAILURE CAN HARM THE OPERATOR OR PATIENT OR DAMAGE THE PRODUCT |
|  | WASTE DISPOSAL |
|  | TYPE B PARTS SUPPLIED |
|  | BATTERY POWER |
|  | INDICATION OF WEIGHT FUNCTIONALITY |
| (O( | INDICATION OF STABLE WEIGHT |
|  | INTERFERENCE COULD OCCUR NEAR THE APPLIANCES |
|  | DUAL INSULATION (CLASS II) |

# 3. SAFETY

ATTENTION!

Before putting the device into use, please read with care the information given in the Operating Instructions. They contain important instructions for installation, proper use and maintenance of the device. The manufacturer shall not be liable for damages arising out of failure to heed the following instructions:

| WARNING! IMPORTANT DIRECTIONS |
| --- |
|  | The assembling of RS300 dynamometer must be effected by skilled operators only and before the use be sure that all hanging sick lift elements, dynamometer and sling are properly assembled, to weight the patient on safety conditions. A faulty assembling of these elements could cause the risk of falling with serious consequences for the patient. |
|  | After assembling phase and before weighing operations, the operator must be sure that: RS300 dynamometer and the sling are on VERTICAL POSITION During weighing operation the hanging sick lift scale is not moving. The operator must ALWAYS follow and hold the sling of the patient during lifting phase, to avoid torsion and sudden movement of both dynamometer and sling that could cause breakages and irreparable damages, with the risk of serious consequences for the patient. DO NOT MOVE the hanging sick lift scale when the patient is leaned. |
|  | During weighing phase by means of RS300 dynamometer, the operator must be sure that the patient remains still, to avoid any instruments torsion and the following uncorrect weighing. |
|  | After weighing phase the operator must ALWAYS follow and hold the sling of the patient, to avoid torsion and sudden movement of both dynamometer and sling that could cause breakages and irreparable damages, with the risk of serious consequences for the patient. |

-When using electrical components under increased safety requirements, always comply with the appropriate regulations.

- Improper installation will render the warranty null and void.

- Ensure the voltage marked on the power supply unit matches your mains power supply.

- This device is designed for use indoors.

- Observe the permissible ambient temperatures for use

- The device meets the requirements for electromagnetic compatibility. Do not exceed the maximum values

specified in the applicable standards.

- Ensure that the patient does not lean against the device risk of falling! If you have any problem, contact

your local  service partner.

-The maximum scale loading must not be exceded. The device beeps when loading exceeds its maximum capacity.

3.1 Manufacturer's guide and declaration – electromagnetic immunity

| Guide and Statement of manufacturer – Electromagnetic emissions |
| --- |
| The electronic scales RS300 model is scheduled for operation in the electromagnetic environment specified below. the customer and the user should ensure that it is used in that environment |
| Emission test | Conformity | Guide to electromagnetic environment |
| RF Emission CISPR11 | Group 1 Class B | The scales RS300 cable model uses RF energy only for its internal function. The refore its RF emissions are very low and probably not cause no interference in electrical equipment. |
| RF Emission CISPR11 | Group 1, Class B | The scales RS300 is suitable for use in all sanitary and hospital buildings, connected to the low voltage public power supply network. |
| Harmonic emission | Class A |  |
| Voltage fluctuations/ flicker emission | Compliant |  |

| Guide and Statement of manufacturer-Electromagnetic Immunity |
| --- |
| The electronic scales RS300 model is scheduled for operation in the electromagnetic environment specified below. the customer and the user should ensure that it is used in that environment. |
| Immunity test | Compliance | Guide to electromagnetic environment |
| Electrostatic discharges (ESD) IEC/EN61000 - 4 -2 | 6kV in contact 8kV in air | The floors should be made of wood, concrete or ceramic. If the floors are covered in synthetic material, the relative humidity should be maximum 30%. |
| Rapid transients/burst IEC/EN61000 - 4 - 4 | 2kV power supply | The power supply should be of the type used typically in commercial or hospital environments. |
| Surge IEC/EN61000 - 4 - 5 | 1kV differential mode | The power supply should be of the type used typically in commercial or hospital environments. |
| Voltage dips, short interruption and voltage variation IEC/EN61000 - 4 - 11 | <5%UT for 0.5 cycle 40%UT for 05 cycle 70%UT for 25 cycle <5%UT for 5 sec | The power supply should be of the type used typically in commercial or hospital environments. Note=UT is the value of the voltage of the feed. |
| Power frequency Magnetic field IEC/EN61000 - 4 - 8 | 3A/m |  |

NOTE: UT is the value of the power supply voltage.

CAUTION

The medical device requires particular electromagnetic compatibility precautions and must be installed and used according to the information provided in the accompanying documents.

| Manufacturer's guide and declaration - Electromagnetic emissions |
| --- |
| The electronic scales RS300 model can be used in the electromagnetic environment specified below. The customer or the user of the seat must ensure that it is used in compliance with the specifications below. |
| Emissions test | Compliance | Guide to electromagnetic environment |
| Radiated/ conducted / emissions IEC/EN61000-4-6 | 3Vrms 150kHz to 80MHz(for appliances that are not life supporting) | Equipment in RF communications and mobile laptops should ot be used near any part of the tooth, including cables, except when they meet the recommended separation distance calculated by the frequency applied to the transmitter. Recommended separation distance d = 1.2 √ P d = 1.2 √ P from 80 MHz to 800 MHz d = 2.3 √ P from 800 MHz to 2.5 GHz where the P and maximum rated power output of the transmitter in Watts (W) according to the manufacturer of the transmitter and the recommended separation distance in meters (m). The intensity of the field of a fixed RF transmitters, as determined by an electromagnetic site, it could be lower than the level of compliance of each range of frequency b. You can check interference in proximity of equipment marked with the following symbol: |
| Radiated/ conducted /emissions IEC/EN61000-4-3 | 3Vrms 150kHz to 80MHz(for appliances that are not life equipment) |  |
| With 80 MHz and 800 MHz is applied the higher frequency range. These guidelines may not apply in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects and people. a) The intensity of the field for fixed transmitters such as base stations for radiotelephone mobile and cordless phones) and land radio mobile, amateur radio equipment, radio transmitters in the AM and FM and TV transmitters cannot be made theoretically and with precision. To establish an electromagnetic environment due to fixed RF transmitters, it should consider the electromagnetic survey of the site. If the field strength measured at the place where you use the instrument exceeds the applicable level of compliance of the above, it should be placed under observation normal operation of the tooth divice. If you notice abnormal performance, it may take additional measures as a different orientation or position of instrument. of frequencies from 150 kHz to 80 MHz should be less than 3 V / m. |

| Recommended separation distances between portable radio equipment unity. |
| --- |
| The scales RS300 model is expected to operate in an electromagnetic environment in which they control the noise radiated RF. The customer or the tooth may be of help to prevent electromagnetic interference by ensuring a minimum distance between mobile communication devices and portable RF (transmitters) and unity, as recommended below, with respect to the maximum power output of equipment radio. |
| Output power rating of the transmitter W | Distance separating the frequency of the transmitter m |
|  | 150 kHz a 80 MHz d = 1,2 √P | 80 MHz a 800 MHz d = 1,2 √P | 800 MHz a 2,5 GHz d = 2,3 √P |
| 0,01 | 0,12 | 0,12 | 0,23 |
| 0,1 | 0,38 | 0,38 | 0,73 |
| 1 | 1,2 | 1,2 | 2,3 |
| 10 | 3,8 | 3,8 | 7,3 |
| 100 | 12 | 12 | 23 |
| For transmitters with maximum rated power output not reported above, the recommended separation distance d in meters (m) can be calculated using the equation applicable to the frequency of the transmitter, where P and the maximum rated power output of the transmitter in Watt (W) according to the manufacturer of the transmitter. Note: With 80 MHz and 800 MHz is applied the higher frequency range. These guidelines may not apply in all situations. Electromagnetic propagation is affected by absorption and reflection from structures, objects and people. |

# 4. FEATURES

| Model | RS 300 |
| --- | --- |
| Factory | Wunder sa.bi. srl- Trezzo s/Adda- Milano Italia |
| Max Division | Max 300kg, e=100g |
| OIML | Class III |
| Unit Weight | Kg (OIML) |
| Display | LCD 25mm con 5 digits |
| Dimension mm | 120 x70x 160 |
| Key & Functions | ON/ZERO/OFF, HOLD, TARE, |
| Power | 6 alkalin batteries mini-stilo code AAA |
| Temperatura operativa | 0°C / 40°C |

|  | ATTENTION! Only the unit kg is approved by OIML (International Organization of Legal Metrology). |
| --- | --- |

# 5. KEY PANNEL

|  | FUNZIONI RS300: HOLD = Indicatore blocco peso NET = peso netto GROSS = peso lordo kg = unità Kg |  |
| --- | --- | --- |

1. ON/OFF/ZERO/TARE：

1. Switch on/off the scale

2. Press this key to zero the scale.  The zero range is +/- 2% of full capacity.

3. Press this key to tare when loads over zero range.

4. In tare mode, press this key again to clear the tarred value when there is no load on the platform.

5. Press this key without release 3 seconds to power off the scale.

2. HOLD:Press this key to lock weight value while weighing. Disable the hold function, press Hold key again.

3. NET ：”Net” function indicator

4. HOLD：”Hold” function indicator

# 6. WEIGHING OPERATION

Before reading detailed instructions on how to use all the weighing functions that are built into your scale, please read the following important guidelines:

-  Always be sure that the display shows `Zero` before use, if it does not then please press the ZERO key.

- The Professional Medical Hoist scale is designed to detect when a stable weight is achieved, your reading should be taken at this point.

|  | ATTENTION! Tilt angle over 3° will cause false reading. During weighing phase by means of RS dynamometer, the operator must be sure that the patient remains still, to avoid any instruments torsion and the following uncorrect weighing. We recommend to not use any lift or Hoist that hanger bar do not swivel 360 degrees. After assembly at the lifter, the dynamometer must not be rotated horizontally. Rotation is only allowed in connection with 360 degree swivel bearings installed in the lift system. |
| --- | --- |
|  | ATTENTION! For measuring the patient ALWAYS SAFE The operator MUST place a pillow under the patient raised WEIGHING THE PATIENT MUST BE STRICTLY OBSERVED INSTRUCTIONS PARAGRAPH NUMBER 3 |

# 7. FUNCTIONS SETTING

Auto-Off Time Adjustment

This feature affects the period of non-operation. Once exceed the setting time, the device will shut off automatically.

Auto off time: 120 sec/180 sec/240 sec/300 sec/off

Buzzer Adjustment

This function affects the buzzer sound ON and OFF by user’s preference.

Buzzer: On/Off

# 8. OPERATION FLOW

A.OFF         BEEP          End

On         Off

120        180      240         300       Off

：

Setting the time for auto time-off - 120/180/240/300/off (seconds)

Buzzer  ON（beep on）/OFF（beep off）

Save the settings

# 9. SETTING UP YOUR SCALE STEP BY STEP

For Example: Set up 180s AUTO-OFF time and Buzzer OFF

Step 1.  Press HOLD for 3 seconds to enter  setup mode.

Step 2.  Press ZERO to enter A.OFF mode when A.OFF is displayed

Step 3.  Press HOLD to select 180 S which means AUTO-OFF TIME is 180 seconds.

Step 4.  Press ZERO to decide and return to A.OFF mode and press HOLD to shift to  BEEP setup mode.

Step 5.  Press ZERO to enter  BEEP setup mode and press HOLD again to select between ON and OFF .

Step 6.  Press ZERO to decide and return to BEEP setup mode.

Step 7.  Press HOLD to select END and press ZERO to finish setting.

# 10. INSTRUCTION FOR BATTERY INSTALLATION

RS uses six AAA size alkaline batteries. Please read the following instruction before using the scale.

1.Find the battery cover at                 2.Remove the battery cover.

the backside of scale

3. Take out the battery case                4.Installing new AAA size batteries.

5. Reinstalling the battery case.              6. Installing the battery cover.

# 11. WARNING INFORMATION

| 1.Low battery indication This warning show that the voltage of battery is very low, please change a new one. |  |
| --- | --- |
| 2. The loading is over limit when power on, please reduce the loading. |  |
| 3. The loading is under limit when power on, please increasing the loading. |  |
| 4. Over load or Counting error– The loading is over limit when power on, please reduce the loading and try again. If the trouble is still exist, please call the service. |  |
| 5. Counting error (too low)– The loading is lower than limit when power on, please increase the loading and try again. If the trouble is still exist, please call the service.。 |  |
| 6.EEPROM error-- The programs of the scale is error, please call the service. |  |

# 12. CLEANING & MAINTENANCE

We would recommend using alcohol based wipes or similar when cleaning the scales.  Please do not use large amounts of water when cleaning the scales as this will cause damage to the scales electronics, you should also refrain from using corrosive liquids or high pressure washers.  Always disconnect the scales from the mains power supply before cleaning.

The machine is sold approved with first metrological test passed (plate with green M). If required, the first periodic test (green label) can be performed by the manufacturer. At the expiry date indicated on the green label, the periodic test must be performed on the machine. Subsequent calibration is always necessary, if one or more safety markings are damaged or the display shows abnormal weights. Always have maintenance performed by qualified personnel. The WUNDER customer technical support service is at your disposal.

ATTENTION!

In some countries, calibration can be performed only by an authorised/qualified agent. Contact your dealer for further

# 13. DISPOSAL OF THE SCALE

EU 2002/96/EC

This product complies with EU Directive 2002/96/EC. The symbol of the barred waste basket shown on the appliance indicates that this product, which, must be treated separately from domestic refuse at the end of its useful lifetime, must be taken to a differentiated collection centre for electrical and electronic devices, or returned to the dealer when a new equivalent device is purchased. The user is responsible for taking the appliance to the appropriate collection facility at the end of its lifetime. Appropriate differentiated collection for submitting the used appliance to environmentally compatible recycling, treatment, and disposal helps to prevent potential negative effects on the environment and on human health, and favours the recycling of the materials of which the product is composed.

For more detailed information regarding the collection systems available, contact your local waste disposal service or the shop where you purchased the scale.

As a consumer, you are obliged by law to return used or discharged batteries. You can deposit your old batteries in the public collection points of your city, or you can deposit them at any retail facility selling batteries of any type that has dedicated collection receptacles. Also in the case of the “demolition” of electrical and electronic appliances, they must be removed and deposited in the dedicated collection receptacles.

NOTE: The following symbols indicate the presence of harmful substances in the batteries:

Pb Pb = batteries that contain lead

Cd Cd = batteries that contain cadmium

Hg Hg = batteries that contain mercury

ATTENTION!

Do not throw the electrical parts or batteries into the household waste. Dispose of the batteries at the collection centres in your area.

# 14. WARRANTY-LIABILITY

WUNDER products are guaranteed under warranty.

The warranty is valid for faults or malfunctions of the instrument due to manufacturing or material defects, and WUNDER reserves the right to repair or replace the article.

In the event that the repairs of the fault or the delivery of the replacement unit do not have a positive outcome, legal measures are to be considered valid. The term of guarantee is two years, taking effect on the date of purchase.

The manufacturer shall not be held liable for damage deriving from any of the following causes: improper or unsuitable storage or use; improper installation or placement in service by the owner or a third party; normal wear and tear from use; changes or modifications; unsuitable or negligent manipulation; excessive use; chemical, electrochemical or electrical interference or interference from humidity; unless such causes may be attributed to negligence on the part of Wunder

# 15. IDENTIFICATION PLATES

# Text boxes

Tasti funzione

On/Off/Zero/Tare

Hold (blocco)

4. NET

5. GROSS

5. GROSS

4. NET

3. HOLD

Press and hold Hold key for 3 seconds.,

Display shows  “Set”.

Notice:

Hold means ”select” “” jump to another selection horizontally.

On/Off/Zero/Taremeans “ enter”  “  ” jump to another selection vertically.